邓兰洪、陈雄立

ICS 77.120.10

H 12

National Standard of the People’s Republic of China

GB/T 20975.25-2020

Replace GB/T 20975.25-2008

Methods for chemical analysis of aluminium and aluminium alloys

Part 25：Determination of elements content

Inductively coupled plasma atomic emission spectrometric method

铝及铝合金化学分析方法

第25部分：元素含量的测定  电感耦合等离子体原子发射光谱法

(English Translation)

Issue date:2020-06-02

Implementation.date:2021-04-01

Issued by General Administration of Quality Supervision, Inspection and

Quarantine of the People’s Republic of China

Standardization Administration of the People’s Republic of China

Foreword

SAC/TC 243 is in charge of this English translation. In case of any doubt about the contents of English translation, the Chinese original shall be considered authoritative.

The GB/T 20975 Methods for chemical analysis of aluminium and aluminium alloys consists of the following 37 parts under the general title:

——Part 1: Determination of mercury content

——Part 2: Determination of arsenic content

——Part 3: Determination of copper content

——Part 4: Determination of iron content

——Part 5: Determination of silicon content

——Part 6: Determination of cadmium content

——Part 7: Determination of manganese content

——Part 8: Determination of zinc content

——Part 9: Determination of lithium content   Flame atomic absorption spectrometry,

——Part 10: Determination of tin content

——Part 11: Determination of lead content

——Part 12: Determination of titanium content,

——Part 13: Determination of vanadium content

——Part 14: Determination of nickel content,

——Part 15: Determination of boron content,

——Part 16: Determination of magnesium content,

——Part 17: Determination of strontium content

——Part 18: Determination of chromium content,

——Part 19: Determination of zirconium content

——Part 20: Determination of gallium content   Butyl rhodamine B spectrophotometry

——Part 21: Determination of calcium content

——Part 22: Determination of beryllium content

——Part 23: Determination of antimony content

——Part 24: Determination of total rare earth content,

——Part 25: Determination of elements content   Inductively coupled plasma atomic emission spectrometric method

——Part 26: Determination of carbon content   Infrared absorption method

——Part 27: Determination of cerium, lanthanum, scandium   Inductively coupled plasma atomic emission spectrometry method

——Part 28: Determination of cobalt content   Flame atomic absorption spectrometry

——Part 29: Determination of molybdenum content   Thiocyanate  spectrophotometry

——Part 30: Determination of hydrogen content   Heating extraction-thermal conductivity method

——Part 31: Determination of phosphorus content   Molybdenum blue spectrophotometry

——Part 32: Determination of bismuth content

——Part 33: Determination of potassium content   Flame atomic absorption spectrometry

——Part 34: Determination of sodium content   Flame atomic absorption spectrometry

——Part 35: Determination of tungsten content   Flame atomic absorption spectrometry

——Part 36: Determination of silver content   Flame atomic absorption spectrometry

——Part 37: Determination of niobium content

This section is Part 25 of GB/T 20975.

This section is drafted in accordance with the rules given in the GB/T 1.1-2009.

This section replaces the GB/T 20975.25-2008(Methods for chemical analysis of aluminium and aluminium alloys Part 25: Inductively coupled plasma atomic emission spectrometry method) in whole. In addition to a number of editorial changes, the following technical deviations have been made with respect to the GB/T 20975.25-2008:

——Add security alerts for standard usage;

——Add 15 elements such as Ag、Ba、 Co、W、Mo、Er、K、Na、Li、Nd、Y、Yb、Hf、P、Sc .(See Chapter1);

——Modify the determination of ranges of 18 elements such as  Fe、Cu、Mn、Ti、V、Sn、Zn、Zr、Ca、Bi、Cr、Ni、Cd、Be、B、Si、Sr、Sb .(See Chapter1, See Chapter 1 in the edition of 2008);

——Add normative reference files (See Chapter 2);

——Add terms and definitions (See Chapter 3);

——Add the analytical reagent and water requirements (See Chapter 5);

——Delete the quality assurance and control (See Chapter 9 in the edition of 2008).

——Add the test report(See Chapter 11).

This standard was proposed by the Nonferrous Industrial Association of China.

This standard was prepared by SAC/TC 243 Chinese Norferrous Metal Technical Standardization Committee.

The previous editions of this section is as follows:

—GB/T 20975.25-2008.

Methods for chemical analysis of aluminium and aluminium alloys

Part 25：Determination of elements content

Inductively coupled plasma atomic emission spectrometric method

Warning -- People using this standard should have formal laboratory experience. This standard does not identify all possible safety issues. Users are responsible for taking appropriate safety and health measures and ensuring compliance with the conditions laid down in the relevant national legislation.

- Scope

This part of GB/T 20975 specifies the determination of elements content in aluminum and aluminum alloy by inductively coupled plasma atomic emission spectrometry which including silicon, iron, copper, gallium, magnesium, manganese, chromium, nickel, zinc, titanium, silver, boron, bismuth, lithium, lead, tin, vanadium, zirconium, barium, beryllium, calcium, cadmium, cobalt, erbium, hafnium, indium, potassium, sodium, molybdenum, neodymium, phosphorus, antimony , scandium ,strontium tungsten, yttrium, ytterbium.

This part is applicable to the determination of silicon, iron, copper, gallium, magnesium, manganese, chromium, nickel, zinc, titanium, silver, boron, bismuth, lithium, lead, tin, vanadium, zirconium, barium, beryllium, calcium, cadmium, cobalt, erbium, hafnium, indium, potassium, sodium, molybdenum, neodymium, phosphorus, antimony, scandium, strontium tungsten, yttrium, ytterbium in aluminum and aluminum alloy. The elements covered in the scope of this method are listed in the Table 1.

Table 1

| Element | Mass concentration range wX/% | Element | Mass concentration range wX/% |
| --- | --- | --- | --- |
| Si | 0.010～13.50 | Be | 0.0001～5.00 |
| Fe | 0.0020～12.00 | Ca | 0.020～10.00 |
| Cu | 0.0005～20.00 | Cd | 0.0020～1.00 |
| Ga | 0.0050～0.050 | Co | 0.0050～12.00 |
| Mg | 0.0010～10.00 | Er | 0.050～11.00 |
| Mn | 0.0010～12.00 | Hf | 0.050～0.50 |
| Cr | 0.0020～5.00 | In | 0.010～0.10 |
| Ni | 0.0020～2.50 | K | 0.0050～1.00 |
| Zn | 0.0010～12.00 | Na | 0.0050～0.50 |
| Ti | 0.0010～15.00 | Mo | 0.0050～2.00 |
| Ag | 0.0050～2.00 | Nd | 20.00～35.00 |
| B | 0.0010～12.00 | P | 0.0050～0.050 |
| Bi | 0.010～11.00 | Sb | 0.010～6.50 |
| Li | 0.0005～10.00 | Sc | 0.50～2.50 |
| Pb | 0.10～1.00 | Sr | 0.0005～10.00 |
| Sn | 0.020～2.00 | W | 0.0050～2.00 |
| V | 0.0010～10.00 | Y | 0.050～11.00 |
| Zr | 0.0020～16.00 | Yb | 10.00～15.00 |
| Ba | 0.0050～0.10 |  |  |

- Normative References

The following referenced documents are essential for the application of this document. For dated references, only the edition cited applies. For undated references, the latest edition of the referenced document (including any amendments) applies.

GB/T 8005.2   Terms of aluminium and aluminium alloys   Part 2: Chemical analysis

GB/T 8170-2008  Rules of rounding off for numerical values ＆expression and judgement of limiting values.

- Terms and definitions

The terms and definitions given in the GB/T 20000.1、GB/T 8005.2 applies.

- Summary of Test method

According to the type of alloy and the content of elements, the following methods are used to decompose the sample:

—Decompose the sample by hydrochloric acid and hydrogen peroxide;

—Decompose the sample by the mixed acid of hydrochloric acid and nitric acid;

—Decompose the sample by sodium hydroxide and hydrogen peroxide;

—Decompose the sample by the mixed acid of hydrochloric acid and nitric acid and hydrofluoric acid;

The solution is nebulizated into an ICP optical emission spectrometer and measured the intensity of the emitted light from the elements. The influence of the matrix shall be calibrated by basing on a very close matrix.

- Reagents

During the analysis, use only reagents of recognized analytical grade and only grade 2 water .

5.1Aluminum, (wAl≥99.999%).

5.2 Hydrogen peroxide（ρ=1.10g/mL）.

5.3 Hydrochloric acid（ρ=1.19g/mL）,Guaranteed reagent.

5.4 Nitric acid（ρ=1.42g/mL），Guaranteed reagent.

5.5 Hydrofluoric acid（ρ=1.14g/mL），Guaranteed reagent.

5.6 Perchloric acid (ρ=1.76g/mL), Guaranteed reagent.

5.7 Hydrochloric acid（1+1）.

5.8 Nitric acid（1+1）.

5.9 Mixed acid: Mix 3 part Hydrochloric acid (5.7) and 1 part Nitric acid (5.8) .

5.10 Sodium hydroxide solution（400g/L）.

5.11Aluminum base solution（20mg/mL）：Weigh 20.00g of aluminum (5.1) and transfer to a 1000mL beaker, then add 600mL hydrochloric acid (5.7) in several portions .Cover with a watch glass and heat gently to complete the dissolution. Add a few drops of hydrogen peroxide (5.2) and boil for a few minutes in order to decompose excess hydrogen peroxide. Cool and transfer the solution into a 1000mL volumetric flask, dilute to the mark with water and mix.

5.12 The standard storage solution for each analytical elements : The certified national reference materials (solutions) is preferred, the solution also can be prepared in Appendix A..

5.13 Standard solution

a) Principles of preparation of multi-element standard solutions: Elements with mutual chemical interference, precipitation and mutual spectral interference should be prepared in groups.

b) Dilute the standard storage solution (5.12) to 100 μg/mL, and maintain the same acidity as the standard storage solution (diluted when used).

c) Dilute the standard storage solution (5.12) to 10 μg/mL, and maintain the same acidity as the standard storage solution (diluted when used).

- Apparatus

Inductively coupled plasma atomic emission spectrometer.

——The Spectrometer Chamber has the function of pumping vacuum or driving gas to ensure the stability of the element test signal with the wavelength below 200nm.

——The resolution of the spectrometer is less than 0.005 nm ( at 200 nm) .

- Sample

The sample shall be in the form of fine drillings or chips with maximum thickness of 1 mm.

- Procedure

8.1 Test portion

Weigh the test samples to the nearest 0.0001 g according to Table 2 and Table 3.

Table 2

| Mass concentration range wX/% | Test portion m0/g | Volume of the sample solution V1/ mL | Volume of the fraction solution V2/mL | Volume of the test solution V3/mL |
| --- | --- | --- | --- | --- |
| - 0.0001～0.05 | - 0.50 | 100 | —— | —— |
| - ＞0.05～1.0 | - 0.25 | 250 | —— | —— |
| - ＞1.0～10.0 | - 0.25 | 250 | 10 | 100 |
| ＞10.0～20.0 | 0.25 | 250 | 5 | 100 |
| ＞20.0～30.0 | 0.25 | 500 | 5 | 100 |

- Table 3

| - Mass concentration range wX/% | Test portion - m0/g | Volume of the sample solution - V1/ mL | Volume of the fraction solution V2/mL | Volume of the test solution V3/mL |
| --- | --- | --- | --- | --- |
| - 0.01～1.0 | - 0.25 | - 250 | - —— | - —— |
| - ＞1.0～10.0 | - 0.25 | - 250 | - 10 | - 100 |
| - ＞10.0～30.0 | - 0.25 | - 500 | - 5 | - 100 |

8.2 Parallel test

Two experiments shall be done independently and calculate the mean mass concentration of each element.

8.3 Blank test

Weigh the same amount of aluminum (5.1) as the sample, and do a blank test with the test portion.

- 8.4 Preparation of the test solution

8.4.1 The first method of dissolution

8.4.1.1 This method is suitable for the determination of ferrum, copper, magnesium, manganese, gallium, titanium, vanadium, indium, tin, chromium, zinc, nickel, cadmium, beryllium, strontium, calcium, barium, potassium, sodium, cobalt, tungsten, molybdenum, erbium, lithium, phosphorus, neodymium, ytterbium, scandium in aluminum and aluminum alloy.

8.4.1.2 Place the test portion (8.1) in Table 2 into a 250 mL glass beaker, add 25 mL of hydrochloric acid (5.7).Cover with a watch-glass and heat gently to decompose dissolution, add a few drops of hydrogen peroxide (5.2) and boil a few minutes in order to complete dissolution .Cool the solution to room temperature, transfer to a volumetric flask of the corresponding volume in Table 2. Dilute to the mark with water and mix. Dilute the test solution according to the calibration curve if necessary.

8.4.1.3 When the content of silicon is more than 0.5% , if there is insoluble residue, the test solution shall be filtered into the corresponding volumetric flask in Table 2, wash the residue with water, at least 3 times. Put the residue and filter paper into a platinum crucible , Make the filter paper to ash ,then burn the residue 5min at 800℃.Cool the residue to room temperature, add 5 mL of hydrofluoric acid (5.5) and a few drops of nitric acid(5.8) , heat gently to decompose dissolution. Add 1 mL of perchloric acid (5.6) and heat gently until the solution is dried. Cool the platinum crucible to temperature, add 5 mL of hydrochloric acid (5.7) and heat gently in order to dissolve the residue, transfer this solution to the test solution and dilute to the mark with water and mix. Dilute the test solution according to the calibration curve if necessary.

8.4.2 The second method of dissolution

8.4.2.1 This method is suitable for the determination of ferrum, copper, magnesium, manganese, gallium, titanium, vanadium, indium, tin, lead, bismuth, antimony, chromium, zinc, nickel, cadmium, beryllium, boron, calcium, silver, barium, cobalt, lithium, molybdenum, neodymium, ytterbium, scandium. in aluminum and aluminum alloy.

8.4.2.2 Place the test portion (8.1) in Table 2 into a 250 mL glass beaker, add 25 mL of mixed acid (5.9).Cover with a watch-glass and heat gently to decompose dissolution, boil a few minutes in order to complete dissolution .Cool the solution to room temperature, transfer to a volumetric flask of the corresponding volume in Table 2. Dilute to the mark with water and mix. Dilute the test solution according to the calibration curve if necessary.

8.4.2.3 When the content of silicon is more than 0.5% , if there is insoluble residue, the test solution shall be filtered into the corresponding volumetric flask in Table 2, wash the residue with water, at least 3 times. Put the residue and filter paper into a platinum crucible , Make the filter paper to ash ,then burn the residue 5min at 800℃.Cool the residue to room temperature, add 5 mL of hydrofluoric acid (5.5) and a few drops of nitric acid(5.8) , heat gently to decompose dissolution. Add 1mL of perchloric acid (5.6) and heat gently until the solution is dried. Cool the platinum crucible to temperature, add 5 mL of hydrochloric acid (5.7) and heat gently in order to dissolve the residue ,transfer this solution to the test solution and dilute to the mark with water and mix. Dilute the test solution according to the calibration curve if necessary.

8.4.3 The third method of dissolution

8.4.3.1 This method is suitable for the determination of silicon, ferrum, copper, magnesium, manganese, titanium, boron, vanadium, chromium, zinc, nickel, zirconium, strontium, tin, antimony, lead, calcium, tungsten, ytterbium, erbium, neodymium and scandium in aluminum and aluminum alloy.

8.4.3.2 Place the test portion (8.1) in Table3 into a 400mL Teflon beaker, add a few water and add 6mL sodium hydroxide solution (5.10). After the severe reaction stop , heat gently to decompose dissolution. Add a few drops of hydrogen peroxide (5.2) and boil a few minutes in order to complete dissolution Then boil several minute until the solution become slimy. Cool and add 30mL water ,heat gently to compose dissolution.

8.4.3.3 Dilute the solution (8.4.3.1) with water to about 100 mL, stir while adding 25 mL nitric acid (5.8) and 25 mL hydrochloric acid (5.7), heat gently to complete dissolution. (add anhydrous sodium sulfite solution to decompose if there is a manganese dioxide brown precipitate) , Cool the solution to room temperature and  transfer the solution into a 250mL volumetric flask, dilute with water to the scale and mix . Dilute the test solution according to the calibration curve if necessary.

8.4.4 The fourth method of dissolution

8.4.4.1 This method is suitable for the determination of zirconium, hafnium, iron, copper, magnesium, manganese, titanium, vanadium, chromium, zinc, nickel, strontium, tin and lead in aluminium and aluminium alloys with silicon content less than 0.50% .

8.4.4.2 Place the test portion (8.1) in Table 2 into a 100 mL Teflon beaker, add 25 mL of mixed acid (5.9) and a few drops of hydrofluoric acid (5.5), Heat gently to decompose dissolution. Cool the solution to room temperature, transfer to a volumetric flask of the corresponding volume in Table 2. Dilute to the mark with water and mix. Dilute the test solution according to the calibration curve if necessary.

8.4.4.3 For easily hydrolyzed elements such as Ti, Sn, Zr, Sb, the test solution should maintain 10% acidity; When making low-content samples, high-purity acid and quartz subboiling distilled water should be used for easily contaminated elements such as Ca, Pb, and B.

8.5 Preparation of the calibration solutions

8.5.1 When the sample is dissolved by methodⅠ, methodⅡ and method Ⅳ, according to the amount of aluminum in the test solution, transfer an appropriate amount of the aluminum base solution (5.11) into a series of 100 mL one-mark volumetric flasks, Add an appropriate amount of the standard storage solution (5.12) or (5.13) and an appropriate hydrochloric acid or nitric acid in order to make the acidity of the calibration solution and test solution basically consistent. Dilute to the mark with water and mix. The test solution without the standard solution is used as the blank solution. The content of the elements to be measured should be within the range of the calibration curve. The number of series of standard solutions is determined by the precision, generally 3 to 5.

8.5.2 When the sample is dissolved by method Ⅲ ,according to the amount of aluminum in the test solution, weigh the appropriate amount of aluminum (5.1) and place into a series of Teflon beaker, then follow the step 8.4.3. Add an appropriate amount of the standard storage solution (5.12) or (5.13) and an appropriate hydrochloric acid or nitric acid in order to make the acidity of the calibration solution and test solution basically consistent. Dilute to the mark with water and mix. The test solution without the standard solution is used as the blank solution. The content of the elements to be measured should be within the range of the calibration curve. The number of series of standard solutions is determined by the precision, generally 3 to 5.

8.5.3 According to the brand of the sample, the corresponding standard sample (national first-class standard sample) can be selected, and a series of calibration solutions can be prepared according to the preparation method of the analytical test solution (8.4).The number of series standard solutions is determined by the precision, generally 3 to 5.

8.6 Determination

Start the inductively couple plasma optical emission spectrometer and let it stabilize before any measurement .Optimize the instrument and select appropriate wavelength . Measure each of the calibration solution two times and calculate the mean intensities. When the linear correlation coefficient of the calibration curve is greater than or equal to 0.9995, Measure the test sample solution two times .

9  Expression of results

Calculate the content of each element in the sample according to Formula (1), it also can be calculated by    the computer:

……………………………(1)

Where

is the concentration of each element in the test solution, in μg/mL；

0	 is the concentration of each element in the blank solution，in μg/mL；

is the volume of the sample solution，in mL；

is the volume of the fraction solution，in mL；

is the volume of the test solution，in mL；

is the mass of test portion，in g。

When the mass concentration of the elements is less than 0.001%, the expression of results retain 1 digit significant figure. When the mass concentration of the elements is between 0.001% and 1.00%, the expression of results retain 2 digits significant figure. When the mass concentration of the elements is more than 1.00%, the expression of results retain 2 decimal place.

10  Precision

10.1 Repeatability

The measured values of two independent test results obtained under repeatability conditions, within the range of average values given below, The absolute difference between the two test results does not exceed the repeatability limit (r), and the case of exceeding the repeatability limit (r) does not exceed 5%, and the repeatability limit (r) is obtained by linear interpolation based on the following data.

Table 4

| wx/% | 0.0005 | 0.0010 | 0.0050 | 0.010 | 0.050 | 0.10 | 0.50 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| r/% | 0.0001 | 0.0003 | 0.0006 | 0.002 | 0.004 | 0.01 | 0.03 |
| wx/% | 1.00 | 5.00 | 10.00 | 15.00 | 20.00 | 25.00 |  |
| r/% | 0.05 | 0.20 | 0.30 | 0.55 | 0.75 | 1.25 |  |

10.2 Reproducibility

The absolute difference between the two independent test results obtained under reproducibility conditions is not greater than the reproducibility limit (R), and the case of exceeding the reproducibility limit (R) is not more than 5%. The reproducibility limit (R) is obtained by linear interpolation based on the following data.

Table 5

| wx/% | 0.0005 | 0.0010 | 0.0050 | 0.010 | 0.050 | 0.10 | 0.50 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| R/% | 0.0002 | 0.0004 | 0.0007 | 0.003 | 0.006 | 0.02 | 0.04 |
| wx/% | 1.00 | 5.00 | 10.00 | 15.00 | 20.00 | 25.00 |  |
| R/% | 0.08 | 0.30 | 0.45 | 0.65 | 1.00 | 1.50 |  |

- Test report

The test report shall include the following information:

- the method used by reference to this document

- all information necessary for the identification of the sample, the laboratory, and the date of analysis or of the test report;

- the results and the form in which they are expressed;

- any unusual features noted during the determination;

- signature of the responsible person.

Annex A

（Informative）

Preparation of standard stock solutions

A.1  Silicon standard stock solution（0.5 mg/mL）：

Accurately weigh 1.0697g of silicon dioxide (G.R) that was previously burned to constant weight at 1000 °C in a platinum crucible, cover a platinum watch glasses, add 5g of a mixture of sodium carbonate and potassium carbonate (1+1), melt to transparent, and cool down. Dissolve the frit with warm water, transfer to a polyethylene beaker, dilute to 700 mL with water, transfer to a 1000 mL volumetric flask, dilute to the mark with water, and mix well. 1 mL of this solution contains 0.5 mg of silicon (stored in a polyethylene bottle).

A.2   Iron standard stock solution (1 mg/mL):

Accurately weigh 1.000g of iron (≥99.99%) into a 400mL beaker, cover the watch glasses, add 40mL hydrochloric acid, slowly heat until completely dissolved, cool, transfer the solution to a 1000mL volumetric flask, and dilute to the mark with water, mix well. 1 mL of this solution contained 1.0 mg of iron.

A.3   Copper standard stock solution (1 mg/mL):

Accurately weigh 1.000g of copper (≥99.99%) into a 400mL beaker, cover the watch glasses, add 10mL nitric acid, slowly heat to completely dissolve, cool, transfer the solution into a 1000mL volumetric flask, and dilute to the mark with water, mix well. 1 mL of this solution contained 1.0 mg of copper.

A.4   Manganese standard stock solution (1 mg/mL):

Accurately weigh 1.000g of manganese (≥99.99%) into a 400mL beaker, cover the watch glasses, add 40mL hydrochloric acid, slowly heat until completely dissolved, cool, transfer the solution into a 1000mL volumetric flask, and dilute to the mark with water, mix well. 1 mL of this solution contained 1.0 mg of manganese.

A.5   Magnesium standard stock solution (1mg/mL):

Accurately weigh 1.000g of magnesium (≥99.99%) into a 400mL beaker, cover the watch glasses, add 40mL hydrochloric acid, slowly heat until completely dissolved, cool, transfer the solution into a 1000mL volumetric flask, and dilute to the mark with water. 1 mL of this solution contained 1.0 mg of magnesium.

A.6   Nickel standard stock solution (1 mg/mL):

Accurately weigh 1.000g of nickel (≥99.99%) in a 400mL beaker, cover the watch glasses, add 40mL of hydrochloric acid, slowly heat to completely dissolve, cool, transfer the solution into a 1000mL volumetric flask, and dilute to the mark with water, mix well. 1 mL of this solution contained 1.0 mg of nickel.

A.7   Zinc standard stock solution (1 mg/mL):

Accurately weigh 1.000g zinc (≥99.99%) into a 400mL beaker, cover the watch glasses, add 40mL hydrochloric acid, slowly heat to completely dissolve, cool, transfer the solution into a 1000mL volumetric flask, and dilute to the mark with water, mix well. 1 mL of this solution contained 1.0 mg of zinc.

A.8   Beryllium standard stock solution (1 mg/mL):

Accurately weigh 1.000 g of beryllium (≥99.99%) into a 400 mL beaker, cover the watch glasses, add 40 mL of hydrochloric acid, slowly heat to completely dissolve, cool, transfer the solution into a 1000 mL volumetric flask, and dilute to the mark with water, mix well. 1 mL of this solution contained 1.0 mg of beryllium.

A.9   Chromium standard stock solution (1mg/mL):

Accurately weigh 2.8285 g of potassium dichromate that was dried at 140 °C in a 400 mL beaker, cover the watch glasses, and add 20 mL of water and 20 mL of hydrochloric acid for dissolution. Add 20mL of hydrogen peroxide dropwise, leave it for 12h to 24h until the yellow color of the solution completely disappears, slowly heat (do not boil) to decompose the excess hydrogen peroxide, cool, transfer the solution to a 1000mL volumetric flask, dilute to the mark with water, and mix. 1 mL of this solution contained 1.0 mg of chromium.

A.10   Titanium standard stock solution (1 mg/mL):

Accurately weigh 1.000g of titanium metal (≥99.99%) into a platinum crucible. After adding a little water, slowly add hydrofluoric acid to dissolve the sample, and then add nitric acid to completely oxidize the low valence titanium. Add 10mL sulfuric acid, shake well, heat and evaporate until the white fumes of sulfuric acid are emitted. After taking off and cooling, transfer the solution into a 1000mL volumetric flask, dilute to the mark with sulfuric acid (5%), and mix well. 1 mL of this solution contained 1.0 mg of titanium.

A.11   Gallium standard stock solution (1 mg/mL):

Accurately weigh 1.000g of gallium (≥99.99%) into a 400mL beaker, cover the watch glasses, add 40mL of hydrochloric acid, add a few drops of nitric acid, slowly heat until completely dissolved, cool, and transfer the solution into a 1000mL volumetric flask, dilute to the mark with 5% hydrochloric acid and mix well. 1 mL of this solution contained 1.0 mg of gallium.

A.12   Vanadium standard stock solution (1 mg/mL):

Accurately weigh 1.7850g of vanadium pentoxide (≥99.99%), which was dried at 110℃ for 1h and cooled to room temperature in a desiccator, put it in a 300mL beaker, and cover the watch glasses. Add 40mL sodium hydroxide, slowly heat until completely dissolved, add 100mL nitric acid after cooling, transfer the solution to a 1000mL volumetric flask, dilute to the mark with water, and mix well. 1 mL of this solution contained 1.0 mg of vanadium.

A.13   Cadmium standard stock solution (1mg/mL):

Accurately weigh 1.000 g of cadmium (≥99.99%), place it in a 400 mL beaker, cover the watch glasses, add 10 mL of water, 30 mL of nitric acid, slowly heat to completely dissolve, cool, and transfer the solution into a 1000 mL volumetric flask. Dilute with water to the mark and mix well. 1 mL of this solution contained 1.0 mg of cadmium.

A.14   Lead standard stock solution (1 mg/mL):

Accurately weigh 1.000g of lead (≥99.99%) into a 300mL beaker, cover the watch glasses, add 10mL nitric acid, slowly heat to completely dissolve, boil for several minutes to remove nitrogen oxides, cool, and transfer the solution to 1000mL In a volumetric flask, dilute to the mark with water and mix well. 1 mL of this solution contained 1.0 mg of lead.

A.15   Tin standard stock solution (0.5 mg/mL):

Accurately weigh 0.5000g of tin (≥99.99%) into a 300mL beaker, cover the watch glasses, add 100mL of hydrochloric acid, slowly heat until completely dissolved, cool, transfer the solution into a 1000mL volumetric flask, and dilute to the mark with water, mix well. 1 mL of this solution contained 0.5 mg of tin.

A.16   Boron standard stock solution (1 mg/mL):

Accurately weigh 5.7174g of boric acid (G.R) dried in a vacuum dryer into a 400mL beaker, cover the watch glasses, add 300mL of water, slightly heat to completely dissolve, cool, transfer the solution into a 1000mL volumetric flask, dilute to the mark with water and mix well. 1 mL of this solution contained 1.0 mg of boron.

A.17   Zirconium standard stock solution (1mg/mL):

Accurately weigh 1.7664g of zirconium oxychloride octahydrate (ZrOCl2•8H2O, G.R) into a 400mL beaker, add 100mL water and 170mL hydrochloric acid  to dissolve, transfer the solution to a 500mL volumetric flask, dilute to the mark with water, and mix well. 1 mL of this solution contained 1.0 mg of zirconium.

A.18   Indium standard stock solution (1 mg/mL):

Accurately weigh 1.0000 g of indium (≥99.99%) into a 400 mL beaker, cover the watch glasses, add 30 mL of hydrochloric acid, heat it on a water bath to completely dissolve, cool, dilute to the mark with water and mix well. 1 mL of this solution contained 1.0 mg of indium.

A.19   Strontium standard stock solution (1 mg/mL):

Accurately weigh 3.0420g of strontium chloride hexahydrate (SrCl2•6H2O) into a 250mL beaker, add water to dissolve, transfer the solution to a 1000mL volumetric flask, dilute to the mark with water, and mix well. 1 mL of this solution contained 1.0 mg of strontium.

A.20   Bismuth standard stock solution (1 mg/mL)

Accurately weigh 1.000 g of bismuth (≥99.99%) into a 400 mL beaker, add 50 mL of nitric acid, cover the watch glasses, slowly heat to completely dissolve, cool, transfer the solution into a 1000mL volumetric flask, and then wash the sides of the beaker with 50mL nitric acid. The washing solution was incorporated into the volumetric flask, diluted with water to the mark and mixed. 1 mL of this solution contained 1.0 mg of bismuth.

A.21   Calcium Standard Stock Solution (1 mg/mL):

Accurately weigh 2.4971g of calcium carbonate (CaCO3, reference reagent) in a 400mL beaker, add 20mL of water, then dropwise add hydrochloric acid  to completely dissolve, and excess 20mL, boil to remove carbon dioxide, cool, transfer the solution into a 1000mL volumetric flask, dilute to the mark with water and mix well. 1 mL of this solution contained 1.0 mg of calcium.

A.22   Antimony standard solution (1mg / mL):

Accurately weigh 1.000g of antimony (≥99.99%) into a 300mL beaker, cover the watch glasses, add 20mL to 30mL sulfuric acid, slowly heat to completely dissolve, cool, and transfer the solution to 1000mL volumetric flask with sulfuric acid, dilute to the mark with sulfuric acid  and mix. 1 mL of this solution contained 1.0 mg of antimony.

A.23   Sodium Standard Stock Solution (1 mg/mL):

Accurately weigh 2.5421g of sodium chloride (NaCl, G.R) in a 400mL beaker, dissolve in the water, transfer the solution into a 1000mL volumetric flask, dilute to the mark with water and mix well. 1 mL of this solution contained 1.0 mg of sodium.

A.24   Potassium Standard Stock Solution (1mg/mL):

Accurately weigh 1.9068g of potassium chloric (KCl, G.R) in a 400mL beaker, dissolve in the water, transfer the solution into a 1000mL volumetric flask, dilute to the mark with water and mix well. 1 mL of this solution contained 1.0 mg of potassium.

A.25   Lithium Standard Stock Solution (1mg/mL):

Accurately weigh 5.3328g of lithium carbonate (LiCO3, G.R) in a 400mL beaker, add 20mL of water, then dropwise add hydrochloric acid to completely dissolve, boil to remove carbon dioxide, cool, transfer the solution into a 1000mL volumetric flask, dilute to the mark with water and mix well. 1 mL of this solution contained 1.0 mg of Lithium.

A.26   Silver standard solution (1mg / mL):

Accurately weigh 1.000g of silver (≥99.99%) into a 300mL beaker, cover the watch glasses, add 25mL nitric acid, slowly heat to completely dissolve, cool, and transfer the solution to 1000mL volumetric flask , dilute to the mark with water and mix. 1 mL of this solution contained 1.0 mg of silver.

A.27   Barium standard solution (1mg / mL):

Accurately weigh 1.7785g of barium chloride (G.R) into a 300mL beaker, dissolve in the boiled water, transfer the solution to 1000mL volumetric flask with water and mix. 1 mL of this solution contained 1.0 mg of barium.

A.28   Cobalt standard solution (1mg / mL):

Accurately weigh 1.000g of cobalt (≥99.99%) into a 300mL beaker, cover the watch glasses, add 50mL nitric acid, slowly heat to completely dissolve, cool, and transfer the solution to 1000mL volumetric flask , dilute to the mark with water and mix. 1 mL of this solution contained 1.0 mg of cobalt.

A.29   Tungsten standard solution (1mg / mL):

Accurately weigh 1.2611g of tungsten trioxide (G.R) into a 300mL beaker, cover the watch glasses, add 20mL to 30mL sodium hydroxide (20%,G.R), slowly heat to completely dissolve, cool, and transfer the solution to 1000mL volumetric flask with water, dilute to the mark with water and mix. 1 mL of this solution contained 1.0 mg of tungsten.

A.30   Molybdenum standard solution (1mg / mL):

Accurately weigh 1.000g of molybdenum trioxide (G.R) into a 300mL beaker, add a few amount sodium hydroxide (20%,G.R), slowly heat to completely dissolve, dilute to 50mL,and transfer the solution to 1000mL volumetric flask with sulfuric acid, dilute to the mark with water and mix. 1 mL of this solution contained 1.0 mg of molybdenum.

A.31   Phosphorus standard solution (1mg / mL):

Accurately weigh 4.2636g of diammonium hydrogen phosphate (G.R) into a 300mL beaker, dissolve with water and transfer the solution to 1000mL volumetric flask with water, dilute to the mark with water and mix. 1 mL of this solution contained 1.0 mg of phosphorus.

A.32   Hafnium standard solution (1mg / mL):

Accurately weigh 1.1793g of hafnium oxide (G.R) into a porcelain crucible, add 4g potassium pyrosulfate, melt transparent at 700 °C for about 10 ~ 15 min, cool, then leach the Frit with 100mL hydrochloric acid (1 + 9) , heat until completely dissolved, transfer to 1000mL volumetric flask, dilute with hydrochloric acid (1 + 9) to scale, shake well. This solution contains 1 mg of Hafnium in 1 mL.

A.33   Yttrium standard solution (1mg / mL):

Accurately weigh 1.2699g of yttrium oxide (G.R) into a 300mL beaker, cover the watch glasses, dissolve in 50mL hydrochloric acid and transfer the solution to 1000mL volumetric flask with water, dilute to the mark with water and mix. 1 mL of this solution contained 1.0 mg of yttrium.

A.34   Neodymium standard solution (1mg / mL):

Accurately weigh 1.1664g of neodymium oxide (G.R) into a 300mL beaker, cover the watch glasses, dissolve in 50mL hydrochloric acid and transfer the solution to 1000mL volumetric flask with water, dilute to the mark with water and mix. 1 mL of this solution contained 1.0 mg of neodymium.

A.35   Erbium standard solution (1mg / mL):

Accurately weigh 1.1435g of Erbium oxide (G.R) into a 300mL beaker, cover the watch glasses, dissolve in 50mL hydrochloric acid and transfer the solution to 1000mL volumetric flask with water, dilute to the mark with water and mix. 1 mL of this solution contained 1.0 mg of erbium.

A.36   Ytterbium standard solution (1mg / mL):

Accurately weigh 1.1387g of ytterbium oxide (G.R) into a 300mL beaker, cover the watch glasses, dissolve in 50mL hydrochloric acid and transfer the solution to 1000mL volumetric flask with water, dilute to the mark with water and mix. 1 mL of this solution contained 1.0 mg of ytterbium.

A.37   Scandium standard solution (1mg / mL):

Accurately weigh 1.5338g of scandium oxide (G.R) into a 300mL beaker, cover the watch glasses, dissolve in 25mL hydrochloric acid and transfer the solution to 1000mL volumetric flask with water, dilute to the mark with water and mix. 1 mL of this solution contained 1.0 mg of scandium.

Annex A

（Informative）

Recommended Line of elements

| Elements | Analytic lines/nm | Elements | Analytic lines /nm |
| --- | --- | --- | --- |
| Si Fe Cu Ga Mg Mn Cr Ni Zn Ti Ag B Bi Li Pb Sn V Zr Ba | 288.158，251.611 259.940，239.562 324.754 294.364 285.213，279.553 259.373，257.610 267.716，283.563 231.604 213.856，206.200 334.941，337.280 328.068 249.678，249.773 223.061 670.784 220.353 189.989 292.402 339.198，349.621 230.424，455.403 | Be Ca Cd Co Er Hf In K Na Mo Nd P Sb Sc Sr W Y Yb | 234.861，313.042 317.933，393.366 228.802 228.616 323.058 399.980 325.609 766.491 589.592 202.030 406.109，401.225 213.618 217.581 361.384，363.075 407.771，346.446 239.709，224.875 224.306，377.433 289.138 |