# ICS 77. 120. 99

H 68

National Standard of the People’s Republic of China

GB/T 15072.13-2008

Replaces GB/T 15072.13-1994

Test methods of precious metal alloys

Determination of tin, cerium and lanthanum contents in silver alloys

Inductively coupled plasma atomic emission spectrometry

贵金属合金化学分析方法

银合金中锡、铈和镧量的测定

电感耦合等离子体原子发射光谱法

(English Translation)

Issue date:2008-03-31                      Implementation date: 2008-09-01

Issued by   General Administration of Quality Supervision, Inspection and

Quarantine of the people's Republic of China

Standardization Administration of the people's Republic of China

Foreword

SAC/TC 243 is in charge of this English translation. In case of any doubt about the contents of English translation, the Chinese original shall be considered authoritative.

This standard is an integrated revision of GB/T 15072-1994( Chemical analysis methods for precious metals and their alloys)(all parts), which is divided into 19 parts:

——GB/T 15072.1-2008 Test methods of precious metal alloys Determination of gold content for gold, platinum and palladium alloys by potentiometric titration with ferrous sulfate;

——GB/T 15072.2-2008 Test methods of precious metal alloys Determination of silver content for silver alloys by potentiometric titration with sodium chloride;

——GB/T 15072.3-2008 Test methods of precious metal alloys Determination of platinum content for gold, platinum and palladium alloys by current titration with potassium permanganate;

——GB/T 15072.4-2008 Test methods of precious metal alloys Determination of palladium content for palladium and silver alloys by butanedione dioxime gravimetry;

——GB/T 15072.5-2008 Test methods of precious metal alloys Determination of silver content for gold and palladium alloys by potentiometric titration with potassium iodide;

——GB/T 15072.6-2008 Test methods of precious metal alloys Determination of iridium content for platinum and palladium alloys by potentiometric titration with ferrous sulfate;

——GB/T 15072.7-2008 Test methods of precious metal alloys Determination of chromium and iron contents for gold alloys by inductively coupled plasma atomic emission spectrometry;

——GB/T 15072.8-2008 Test methods of precious metal alloys Determination of copper content for gold, palladium and silver alloys by EDTA complexometric back titration with thiourea precipitation;

——GB/T 15072.9-2008 Test methods of precious metal alloys Determination of indium content for gold alloys by EDTA complexometric back titration;

——GB/T 15072.10-2008 Test methods of precious metal alloys Determination of nickel content for gold alloys by EDTA complexometric back titration;

——GB/T 15072.11-2008 Test methods of precious metal alloys Determination of gadolinium and beryllium contents for gold alloys by inductively coupled plasma atomic emission spectrometry;

——GB/T 15072.12-2008 Test methods of precious metal alloys Determination of vanadium content for silver alloys by hydrogen peroxide spectrophotometry;

——GB/T 15072.13-2008 Test methods of precious metal alloys Determination of tin, cerium and lanthanum contents for silver alloys by inductively coupled plasma atomic emission spectrometry;

——GB/T 15072.14-2008 Test methods of precious metal alloys Determination of aluminium and nickel contents for silver alloys by inductively coupled plasma atomic emission spectrometry;

——GB/T 15072.15-2008 Test methods of precious metal alloys Determination of nickel, zinc and manganese contents for gold, silver and palladium alloys by inductively coupled plasma atomic emission spectrometry;

——GB/T 15072.16-2008 Test methods of precious metal alloys Determination of copper and manganese contents for gold alloys by inductively coupled plasma atomic emission spectrometry;

——GB/T 15072.17-2008 Test methods of precious metal alloys Determination of tungsten content for platinum alloys by tungsten trioxide gravimetry;

——GB/T 15072.18-2008 Test methods of precious metal alloys Determination of zirconium and gallium contents for gold alloys by inductively coupled plasma atomic emission spectrometry;

——GB/T 15072.19-2008 Test methods of precious metal alloys Determination of vanadium and magnesium contents for silver alloys by inductively coupled plasma atomic emission spectrometry

This part is the thirteenth of GB/T 15072-2008.

This part replaces GB/T 15072.13-1994 (Methods for chemical analysis of precious metals and their alloys Silver alloys Determination of cerium content) in whole.

The following deviations have been made with respect to the GB/T 15072.13-1994(the previous edition):

——The standard tilte of Methods for chemical analysis of precious metals and their alloys Silver alloys Determination of cerium content replaces Chemical analysis methods for precious metals and their alloys Determination of cerium content in silver alloys;

——Inductively coupled plasma atomic emission spectrometry replaces spectrophotometry;

——The scope of the method is extended from AgCe0.5 alloy of the previous standard to the series of AgCe, AgSnCeLa, AgCuSnCe, AgCuNiCe alloys, of which AgSnCeLa, AgCuSnCe, AgCuNiCe are new;

——The range of cerium content(mass fraction) is extended from 0.3% ~ 0.7% to 0. 2% ~ 2%;

——The simultaneous determination of lanthanum and tin was added, and the range (mass fraction) is 0.2%~2%;

—— Repeatability and relative deviation replaces the allowed deviation of previous standard.

Annex A is informative.

This part was proposed by Nonferrous Metals Industry Association of China.

This part was prepared by SAC/TC 243 Chinese Nonferrous Metal Standardization Technical Committee.

The previous edition of this part is as follows:

——GB/T 15072.13-1994.

Test methods of precious metal alloys—Determination of tin, cerium and lanthanum contents in silver alloys—Inductively coupled plasma atomic emission spectrometry

1   Scope

This part specifies a method for the determination of tin, cerium and lanthanum contents in silver alloys.

The part is applicable to the determination of tin, cerium and lanthanum contents between 0.2% and 2 % (mass fraction) in AgCe, AgSnCeLa, AgCuSnCe, AgCuNiCe alloys.

2   Normative references

The following normative documents contain provision which, through reference in this test, constitute provisions of this part. For dated references, subsequent amendments (excluding corrections), or revisions, of any of these publications do not apply to this part. However parties to agreements based on this part are encouraged to investigate the possibility of applying the most resent editions of the normative documents indicated below. For undated references, the latest edition of the normative document referred to applies.

YS/T 371 Methods for chemical analysis of precious metals alloys General rules and regulations.

3   Method abstract

The test portion shall be dissolved in nitric acid. Sulfuric acid is added in the case of the sample contained tin for preventing hydrolysis. Silver is precipitated and separated with hydrochloric acid. The cerium, lanthanum and tin contents can be determined by inductively coupled plasma atomic emission spectrometry (ICP-AES) and calculated as mass fraction.

4   Reagents and materials

Unless otherwise stated, use only reagents and materials as specified in YS/T 371.

4.1   Argon, purity≥99.99%(mass fraction).

4.2   Nitric acid, ρ about 1,42g/mL.

4.3   Hydrochloric acid, ρ about 1,19 g/mL.

4.4   Sulfuric acid, ρ about 1,84g/mL.

4.5   Hydrogen peroxide, purity 30% (mass fraction).

4.6   Nitric acid, heat gently to evaporate nitrogen dioxide, add one portion of nitric acid to one portion of water.

4.7   Hydrochloric acid, diluted 1+9.

4.8   Cerium standard solution

Weighing 0.1228g of cerium oxide (purity≥99.95%(mass fraction)) and transfer into a 200 mL beaker. Add 10 mL of nitric acid (4.6) and 3 mL of hydrogen peroxide (4.5). Heating the solution until test portion completely dissolved. It should be covered with a watch-glass and add hydrogen peroxide several times until cerium dioxide dissolves completely. Cool to room temperature. The solution shall be transferred into a 100 mL one-mark volumetric flask with hydrochloric acid (4.7) and dilute to the mark and mix. 1 mL of this solution contains 1mg of cerium.

4.9   Lanthanum standard solution

Weighing 0.1173g of lanthanum oxide (purity≥99.95%(mass fraction)) and transfer into a 200 mL beaker. Add 10 mL of hydrochloric acid (4.7). It should be covered with a watch-glass, Heat the solution until test portion completely dissolved. Cool to room temperature. The solution shall be transferred into a 100 mL one-mark volumetric flask with hydrochloric acid (4.7) and dilute to the mark and mix. 1 mL of this solution contains 1mg of lanthanum.

4.10   Tin standard solution

Weighing 0,1000g of tin metal (purity≥99.95%(mass fraction)) and transfer into a 200 mL beaker. Add 10 mL of hydrochloric acid (4.7). It should be covered with a watch-glass, Heat the solution until test portion completely dissolved. Cool to room temperature. The solution shall be transferred into a 100 mL one-mark volumetric flask with hydrochloric acid (4.7) and dilute to the mark and mix. 1 mL of this solution contains 1mg of tin.

5   Apparatus

Inductively coupled plasma atomic emission spectrometry.

The instrument used will be satisfactory if, after optimizing in accordance with the manufacturer’s instructions, it meets the performance criteria as the following.

5.1   Light source

Argon plasma light source, the maximum output power of the generator is not less than 1.3KW.

5.2   Resolution

The optical resolution is not more than 0.010nm at about 200nm and not more than 0.020nm at about 400nm.

5.3   Instrument stability

The stability of the instrument within one hour is less than 2.0%, expressed as RSD.

5.4   Detection limit of spectrometer

The detection limits of cerium, lanthanum and tin in blank solution are not more than 0.05mg/l.

6   Sample

The sample in the form of drillings or flakes shall be processed into chips, then cleaned with acetone to degrease, rinsed with water, dried and mixed well.

7   Analysis procedure

7.1   Test portion

Weigh 0.1g of test portion, accurate to 0.1mg,

Two independent determinations should be required, and average the results.

7.2   Blank test

Blank test shall be done along with the test portion.

7.3   Analysis procedure

7.3.1   Preparation of the test solution

7.3.1.1   Preparation of the test solution in the case of the alloy sample not containing tin

The test portion (7.1) should be placed into a 50 mL beaker. Add 3 mL of nitric acid (4.2). Cover with a watch-glass and heating in low temperature until the solution was completely dissolved. Add 10 mL of hydrochloric acid (4.7), slight boil for 20 min to silver precipitated completely. Transfer the precipitate and solution into 100 mL one-mark volumetric flask. Dilute to the mark with hydrochloric acid (4.7) and mix. It should be reserved for more than 30 min, and take the measurement until the solution clear.

7.3.1.2   Preparation of the test solution in the case of the alloy sample containing tin

The test portion (7.1) should be placed into a 50 mL beaker. Add 1 mL of sulfuric acid (4.4) and 3 mL of nitric acid (4.2). Cover with a watch-glass and heating in low temperature until the solution was completely dissolved. Rinse the watch-glass and the inside of beaker with a small quantity of water. Add 10 mL of hydrochloric acid (4.7), slight boil for 20 min to silver precipitated completely. Transfer the precipitate and solution into 100 mL one-mark volumetric flask. Dilute to the mark with hydrochloric acid (4.7) and mix. It should be for more than 30 min, and take the measurement until the solution clear.

7.3.2   Preparation of the calibration solution

Transfer appropriate volume of cerium, lanthanum and tin standard solution (4.8～4.10) respectively into a series of 100 mL one-mark volumetric flasks, which is diluted step by step. Dilute to the mark with hydrochloric acid (4.7) and mix. Such, obtain the calibration solution shown in Table 1.

Table 1                                         in microgram per millilitre

| Calibration solution label | 0# | 1# | 2# | 3# | 4# |
| --- | --- | --- | --- | --- | --- |
| Concentration of tin | 0.00 | 0.50 | 5.00 | 10.00 | 25.00 |
| Concentration of cerium | 0.00 | 0.50 | 5.00 | 10.00 | 25.00 |
| Concentration of lanthanum | 0.00 | 0.50 | 5.00 | 10.00 | 25.00 |

7.4   Determination

7.4.1   Condition of determination

The condition of determination is showed in Annex A.

7.4.2   Analytical line

Recommended wavelengths are given in Table 2.

Table 2

| Element | Wavelength/nm |
| --- | --- |
| Ce | 418.660 |
| La | 379.478 |
| Sn | 283.998 |

7.4.3   Preparation of working curve

Input mixed differential standard solution (7.3.2) into ICP-AES to measure cerium, lanthanum and tin in accordance with the prepared procedure. Establish the calibration curve using the concentrations of elements to be measured on the x-axis and the corresponding wavelength intensities on the y-axis. The correlation coefficient of calibration curve shall be higher than 0.9999.

7.4.4   Measurement of the test solution

Input the test sample solution (7.3.1.1, 7.3.1.2) to ICP-AES in accordance with the prepared procedure.

8   Expression of results

The mass fraction of elements to be measured, expressed as a percentage, is given by Formula (1):

wX=ρx⋅Vm0×106×100                                               (1)

In the formula:

ρx  is the concentration of cerium, lanthanum and tin obtained from the working curve, in µg/mL;

V    is the volume of the test solution, in mL;

m0   is the mass of the test portion, in g.

Calculate the mass fraction to the third decimal place in the case of a mass fraction more than 1%, the second decimal place in the case of a mass fraction less than 1%.

9   Precision

9.1   Repeatability

Within the mean values given below, the absolute difference in results of two parallel determinations under repeatability conditions shall not be greater than repeatability limit. If the difference exceeds the limit, it shall be not more than 5%.The repeatability limit is obtained by linear interpolation according to the data in Table 3.

Table 3                                                      in %

| Mass fraction | 0.20 | 0.71 | 2.00 |
| --- | --- | --- | --- |
| repeatability limit of cerium | 0.01 | 0.019 | 0.027 |
| Repeatability limit of lanthanum | 0.01 | 0.020 | 0.030 |
| repeatability limit of tin | 0.08 | 0.023 | 0.034 |

9.2   Tolerance

Table 4                                                       in %

| Mass fraction of tin, cerium and lanthanum | relative tolerance |
| --- | --- |
| ≥0.20～0.50 | 5.0 |
| ≥0.50～1.00 | 4.0 |
| ≥1.00～2.00 | 3.0 |

Annex A

(informative)

Working conditions of the apparatus

Optimize the instrument in accordance with the operating manual. Set up the parameters refer to as follows:

A.1 High Radio Frequency power of the generator is 1.2KW.

A.2 Argon flow-rate: of which cooling gas is 15 l/min, protection gas is 0.8l/min, and carrier gas is 0.3 l/min.

A.3 Vertical observation height is 15 mm.

A.4 Integration time is 5 s.

A.5 Flow velocity of injection pump is 15mL/min.