GB/T 12690.—

GB/T 18114.3-2025


[image: image1.wmf]w


GB/T 18114.3—2025
                                                     Replaces GB/T 18114.3—2010



Foreword

SAC/TC 229 is in charge of this English translation. In case of any doubt about the contents of the English translation, the Chinese original shall be considered authoritative.

This document is drafted in accordance with the rules given in the GB/T 1.1—2020 Directives for standardization — Part 1: Rules for the structure and drafting of standardizing documents.
This document is Part 3 of GB/T 18114 Chemical analysis methods for rare earth concentrates. The following parts have been released for GB/T 18114:
—Part 1: Determination of total rare earth oxide contents-Gravimetry;

—Part 2: Determination of oxide thorium content;

—Part 3: Determination of calcium oxide content;

—Part4: Determination of niobium oxide, zirconium oxide and titanium oxide contents-Inductively coupled plasma atomic emission spectrometry;

—Part 5: Determination of aluminium oxide content-Inductively coupled plasma atomic emission spectrometry;

—Part 6: Determination of silicon oxide content;

—Part 7: Determination of iron oxide content-Potassium dichromate titration;

—Part 8: Determination of rare earth oxide contents and distribution;

—Part 9: Determination of phosphorus pentoxide content;

—Part 10: Determination of water content - Gravimetry;

—Part 11: Determination of fluorine content.

This document replaces GB/T 18114.3-2010 Chemical Analysis Methods of Rare Earth Concentrates — Part 3: Determination of calcium oxide content. Compared with GB/T 18114.3-2010, except for structural adjustments and editorial changes, the main technical changes are as follows:

a) Changed the scope of application of the document (see Clause 1, Clause 1 of the 2010 edition);
b) The test portion of the samples in the Flame Atomic Absorption Spectrometry (method 1) was changed from 0.1 g to 0.3 g (see 4.5.1, and 6.1 of the 2010 edition);

c) In the Flame Atomic Absorption Spectrometry (method 1), the flux for the test samples is changed from "2 g sodium hydroxide + 2 g sodium peroxide" to "2 g sodium hydroxide + 2 g sodium peroxide + 2 g sodium carbonate" (see 4.5.4.1, 6.4.1 of the 2010 edition);

d) In Flame Atomic Absorption Spectrometry (method 1), the washing solution has been changed from "sodium hydroxide washing solution (20 g/L)" to "sodium carbonate washing solution (10 g/L)" (see 4.2.7, and 3.7 of the 2010 edition);

e) Changed the aliquot volume of the test solutionfor sampleswith different calcium oxide content in Flame Atomic Absorption Spectrometry(method1)(see Table 1,Table 1 of the 2010 edition);

f) Changed the range of the standard addition calibration curve for the Flame Atomic Absorption Spectrometry method (method 1)(see 4.5.5, and 6.4.2 of 2010 edition);

g) The term " allowable deviation" for precision has been changed to "reproducibility" (see 4.7.3 and 5.7.3, 8.2 and 16.2 of the 2010 edition);

h) The preparation method of the ammonium chloride wash solution and the calibration method for the EDTA standard titration solution in EDTA titration method (method 2) have been changed (see 5.2.15, 5.2.18, and 12.17 and 12.20.2 of 2010 edition);

i) Changed the test portion mass and the aliquot volume in the EDTA titration method (method2)(see Table 4, Table 4 of 2010 edition);

j) Changed the requirements for filter paper and washing solution, the amount of ammonium chloride added, and the amount of magnesium sulfate added in the EDTA titration method(method 2)(see 5.4.4.2、5.4.4.3and5.4.5,14.4.2、14.4.3 and 14.4.4 of 2010 edition);

k) Added the method of adjusting the pH value of the solution using malachite green indicator for the EDTA titration method(method 2) (see 5.4.5);

l) Added the determination requirements for Inductively Coupled Plasma Atomic Emission Spectrometry  method(method 3)(see chapter 6).

Attention is drawn to the possibility that some of the elements of this document may be the subject of patent rights. The issuing body of this document shall not be held responsible for identifying any or all such patent rights.

This document was proposed and prepared by SAC/TC 229 (China Rare Earth Standardization Technical Committee).
This document was issued in 2000 as first edition and first revised in 2010; this is the second revision.

Introduction

Rare earth product with rare earth content up to the requirement of smelting after rare earth ore is beneficiated and enriched. It has not any change in the natural distribution of rare earth elements. Rare earth concentrates include bastnaesite concentrate, monazite concentrate, mixed concentrate of bastnaesite and monazite, bastnaesite-(La) concentrate and so on. Rare earth has significant physical properties in light, electricity and magnetism. Therefore, it can be combined with other materials to form a wide variety of new raw materials with different properties and significantly enhance the quality and performance of other products. Chemical composition are important assessment indexes for rare earth concentrate. By defining the scope of application, specifying reagents, materials, apparatus, and procedures, and by generating a substantial amount of experimental and precision verification data, scientific and accurate standard methods for chemical analysis are established. These standards enhance the consistency and comparability of data among different laboratories, provide rigorous and standardized technical means for quality verification of rare earth concentrate products, and promote the production and trade of these products.
The standards series GB/T 18114, Chemical Analysis Methods for Rare Earth Concentrates, focuses mainly on total rare earth oxide contents, fifteen rare earth oxide relative contents and non-rare earth elements in rare earth concentrates. It comprises 11 parts.

—Part 1: Determination of total rare earth oxide contents—Gravimetry. The aim is to establish a gravimetric method for determining total rare earth oxide contents in rare earth concentrates;

—Part 2: Determination of oxide thorium content. The aim is to establish inductively coupled plasma atomic emission spectrometry and inductively coupled plasma mass spectrometry for determining oxide thorium content in rare earth concentrates;

—Part 3: Determination of calcium oxide content. The aim is to establish flame atomic absorption spectrometry, EDTA titrimetric method and inductively coupled plasma atomic emission spectrometry for determining calcium oxide content in rare earth concentrates;

—Part 4: Determination of niobium oxide, zirconium oxide and titanium oxide contents—Inductively coupled plasma atomic emission spectrometry. The aim is to establish inductively coupled plasma atomic emission spectrometry for determining niobium oxide, zirconium oxide and titanium oxide contents in rare earth concentrates;

—Part 5: Determination of aluminum oxide content—Inductively coupled plasma atomic emission spectrometry. The aim is to establish inductively coupled plasma atomic emission spectrometry for determining aluminum oxide content in rare earth concentrates;

—Part 6: Determination of silicon oxide content. The aim is to establish molybdenum blue spectrophotometry and inductively coupled plasma atomic emission spectrometry for determining silicon oxide content in rare earth concentrates;

—Part 7: Determination of iron oxide content—Potassium dichromate titration. The aim is to establish potassium dichromate titration for determining iron oxide content in rare earth concentrates;

—Part 8: Determination of rare earth oxide contents and distribution. The aim is to establish inductively coupled plasma atomic emission spectrometry for determining rare earth oxide relative contents in rare earth concentrates and X-ray fluorescence spectrometry for determining rare earth oxide contents in rare earth concentrates;

—Part 9: Determination of phosphorus pentoxide content. The aim is to establish phosphorus-bismuth-molybdenum blue spectrophotometry and inductively coupled plasma atomic emission spectrometry for determining phosphorus pentoxide content in rare earth concentrates;

—Part 10: Determination of water content—Gravimetry. The aim is to establish a gravimetric method for determining water content in rare earth concentrates;

—Part 11: Determination of fluorine content. The aim is to establish EDTA titration, alizarin spectrophotometry and fluorine ion selective electrode method for determining fluorine content in rare earth concentrates.

All methods in this document were verified by interlaboratory testing. The term "tolerance" for precision has been replaced by "reproducibility". Repeatability and reproducibility limits for samples with high, medium, and low calcium oxide content are provided, making the precision requirements of the method more comprehensive.
The GB/T18114.3—2000 standard was revised in 2010 for the first time after its initial release and implementation. In order to better adapt to changes in the rare earth industry and to improve the analytical technology and standard system for rare earth concentrates, based on the new industry and market changes and demands, the original methods Flame Atomic Absorption Spectrometry and EDTA titration method are improved. On the basis of this, the Inductively Coupled Plasma Atomic Emission Spectrometry (method 3) is added so that the standard covers the existing focused indicators of calcium oxide content in the current rare earth concentrate products. This provides a fast and accurate method standard for the determination of the chemical composition of rare earth concentrates, and has good operability.

Chemical analysis methods of rare earth concentrates—Part 3: Determination of calcium oxide content

1 Scope

This document describes methods for the determination of calcium oxide content in rare earth concentrates.

This document is applicable to the determination of calcium oxide content in rare earth concentrates.

This document contains three methods in total: Flame Atomic Absorption Spectrometry (method 1), EDTA Titration (method 2), and Inductively Coupled Plasma Atomic Emission Spectrometry (method 3). Determination range (mass fraction) of method 1: 0.50% to 3.00%, Determination range (mass fraction) of method 2: 2.00% to 20.00%, Determination range (mass fraction) of method 3: 0.50% to 20.00%.

When the determination range of 0.50% to 5.00% overlaps, method 3 is used as arbitration method. When the determination range of >5.00% to 20.00% overlaps, method 2 is used as arbitration method.

2 Normative references

The following referenced documents are indispensable for the application of this document. For dated references, only the edition cited applies. For undated references, the latest edition of the reference document (including any amendments) applies.

—GB/T 601 Chemical reagent—Preparations of reference titration solutions.

—GB/T 6379.2 Accuracy (trueness and precision) of measurement methods and results--Part 2: Basic method for the determination of repeatability and reproducibility of a standard measurement method.

—GB/T 6682 Water for analytical laboratory use - Specification and test methods.
—GB/T 8170 Rules of rounding off for numerical values & expression and judgement of limiting values.
—JJG 694 Atomic absorption spectrophotometers.
—JJG 768 Emission spectrometer.

3 Terms and Definitions

No term or definition need to be defined in this document.

4 Flame Atomic Absorption Spectrometry (method 1)
4.1 Principle

The samples are fused with sodium hydroxide, sodium carbonate and sodium peroxide, and extracted by hot water. The sodium carbonate washing solution is used to wash and filter out aluminum, silicon, phosphorus, etc. After filtration, the precipitate and filter paper are dissolved with hydrochloric acid and hydrogen peroxide. The filter paper is crushed into pulp. The solution and the pulp are bottled for dry filtration. During the determination, lanthanum and EDTA are added to eliminate the interference of coexisting ions. In a hydrochloric acid medium, the standard addition method is employed to determine the calcium oxide content using an air-acetylene flame in an atomic absorption spectrometer.

4.2 Reagents and materials 

Unless otherwise stated, only reagents confirmed to be of analytical grade or above and secondary water complying with GB/T 6682 are used in the analysis. All of the liquid reagents shall be stored in plastic bottles, certified standard solutions are preferred.

4.2.1 Sodium hydroxide.
4.2.2 Anhydrous sodium carbonate. 
4.2.3 Sodium peroxide.
4.2.4 Hydrochloric acid (ρ= 1.19 g/mL).

4.2.5 Hydrogen peroxide [mass fraction 
[image: image12.png]


(H₂O₂)≥30%].

4.2.6 Hydrochloric acid(1+1).

4.2.7 Sodium carbonate solution (10 g/L).

4.2.8 Lanthanum chloride solution (
[image: image2.wmf]）

（

La

w

 = 1 g/L).

4.2.9 Disodium ethylenediaminetetraacetate (EDTA) solution (40 g/L).

4.2.10 Calcium oxide standard stock solution: Weigh 0.1785 g of calcium carbonate (with a purity of ≥ 99.99%) (previously dried at 105℃-110℃until constant weight), add 10 mL water, then add 10 mL hydrochloric acid (4.2.4) for dissolution. Transfer it into a 100 mL volumetric flask, dilute the solution with water to the mark, and mix it well. This solution contains 1 mg of calcium oxide per milliliter.

4.2.11 Calcium oxide standard solution: Take 25.00 mL of the calcium oxide standard stock solution (4.2.10) and transfer it into a 500 mL volumetric flask. Add 50 mL of hydrochloric acid (4.2.4), then dilute the solution with water to the mark and mix it well. This solution contains 50 μg of calcium oxide per milliliter.

4.2.12 Acetylene gas: Volume fraction ≥ 99.99%.

4.3 Apparatus

4.3.1 Atomic absorption spectrometer, equipped with an air-acetylene burner.

Under the optimal work conditions, all of atomic absorption spectrometer that can fulfill the following indicators can apply:

— Characteristic concentration: In a solution that is consistent with the matrix of the sample being tested, the characteristic concentration of calcium oxide is not more than 0.15 μg/mL.
— Precision: Measure the absorbance 10 times with the highest concentration standard solution, and the standard deviation shall not exceed 1.0% of the average absorbance value. Measure the absorbance 10 times with the lowest concentration standard solution (not the zero-concentration solution), and the standard deviation shall not exceed 0.5% of the average absorbance value of the highest concentration standard solution.

— Linear working curve: Divide the working curve into five equal sections based on concentration. The ratio of the absorbance difference in the highest section to that in the lowest section should be no less than 0.7.

— Compliant with the provisions of JJG694.

4.3.2 Electronic balance: Resolution 0.1 mg.

4.4 Sample

4.4.1 The particle size of the sample should not exceed 0.074 mm.

4.4.2 The sample is dried in 105℃ to 110 ℃ for 2 hours, then placed in a desiccator and cooled to room temperature.

4.5 Test Procedures

4.5.1 Test portion

Weigh 0.30 g of the sample (4.4), accurate to 0.0001 g.

4.5.2 Parallel test

Take parallel test portions for determination and calculate the mean result.

4.5.3 Blank test 

Carry out a blank test along with the sample.

4.5.4 Preparation for analytical test solution

4.5.4.1 Transfer the sample (4.5.1) to a 30 mL nickel crucible that contains 2.0 g of dehydrated sodium hydroxide (4.2.1) and 2.0 g of anhydrous sodium carbonate (4.2.2). Cover it with 2.0 g of sodium peroxide (4.2.3). Heat it on an electric furnace to remove the moisture, then cover the crucible and place it in a 750℃ muffle furnace to melt until it turns to a bright red color and maintain for 10 minutes (take it out and shake once during the process). Then remove it and cool it.

4.5.4.2 Place the crucible and its cover in a 300 mL beaker, and add 100 mL of hot water to leach the melt. After the intense reaction stops, rinse the crucible and its outer wall with water, add 2 mL of hydrochloric acid (4.2.6) to wash the crucible, then wash it with water and remove the crucible and its cover. Control the volume to approximately 150 mL. Boil the solution for 2 minutes, then cool slightly. Filter it using medium-speed quantitative filter paper, wash the beaker twice or three times with sodium carbonate washing solution (4.2.7), and wash the precipitate five or six times. Put the precipitate together with the filter paper back into the original beaker, add 10 mL of hydrochloric acid (4.2.6) and 1 mL of hydrogen peroxide (4.2.5), heat to dissolve the precipitate, destroy the filter paper, and transfer the solution along with the pulp into a 100 mL volumetric flask, dilute to the mark with water, mix it well, and dry filter. 

4.5.5 Determination

4.5.5.1 According to Table 1, take 5 analytical test solutions (4.5.4.2) and transfer them into a group of 50 mL volumetric flask. Add 5 mL hydrochloric acid (4.2.6), 2.5 mL lanthanum chloride solution (4.2.8), 5 mL Disodium ethylenediaminetetraacetate (EDTA) solution (4.2.9), and 0 mL, 0.50 mL, 1.00 mL, 2.00 mL, 3.00 mL calcium oxide standard solution (4.2.11) to the 5 volumetric flask respectively. Dilute them to the mark with water and mix well.

4.5.5.2 Use an air-acetylene flame and water as the zero reference, measure the absorbance of calcium oxide at a wavelength of 422.7 nm on the atomic absorption spectrometer. Taking the mass concentration of calcium oxide as the horizontal coordinate and the absorbance as the vertical coordinate, a standard addition curve is drawn. The mass concentration of calcium oxide in the test solution was determined using the extrapolation method. 
Table 1 Transfer volume of sample solution for method 1

	Mass fraction of calcium oxide %
	Transfer volume of sample solution mL

	0.50~1.00
	5.00

	>1.00~2.00
	2.00

	>2.00~3.00
	1.00


4.6 Test data processing

The calcium oxide content is calculated as a mass fraction [image: image3.wmf]w

(CaO) using to formula (1): 


[image: image4.wmf]%

100

10

-

2

1

6

1

0

0

a

´

´

´

´

´

=

-

V

m

V

V

O

C

）

（

）

（

r

r

w

…………………………………… (1)
Where,

ρ is the mass concentration of calcium oxide in the analytical test solution, in micrograms per milliliter ( μg/mL);
ρ0 is the mass concentration of calcium oxide in the blank test solution, in micrograms per milliliter (μg/mL);
V0 is the total volume of test solution, in milliliter (mL);
V1 is the volume of test solution to be tested, in milliliter (mL);
m1 is the mass of test material, in gram (g);
V2 is the volume of test solution taken separately, in milliliter (mL).
The absolute difference between two parallel determinations should not exceed the repeatability limit (r) specified in Table 2. The average value is reported as the final result. Results are retained to two decimal places and rounded conforming to the provisions of GB/T 8170.

4.7 Precision

4.7.1 Precision raw data and statistical analysis

The precision data were based on the collaborative trial statistics and calculations from 7 laboratories in 2024, using 3 different levels of samples. Each laboratory independently determined the content of calcium oxide under repeatability conditions, with 11 determinations for each level. The experimental data were statistically analyzed according to GB/T 6379.2.

4.7.2 Repeatability

Under repeatability conditions, the absolute difference between two independent test results should not exceed the repeatability limit (r), and the occurrence of exceeding the repeatability limit (r) should not exceed 5%. The repeatability limit (r) is determined using linear interpolation or extrapolation based on the data in Table 2.

Table 2 The repeatability limit of method 1(r)

	Mass fraction %
	Repeatability limit (r) %

	0.73
	0.07

	1.43
	0.12

	3.00
	0.33

	Note: The repeatability limit (r) is 2.8×Sr ,where Sr  is the standard deviation of repeatability.


4.7.3 Reproducibility

Under reproducibility conditions, the absolute difference between two independent test results should not exceed the reproducibility limit (R), and the occurrence of exceeding the reproducibility limit (R) should not exceed 5%. The reproducibility limit (R) is determined by linear interpolation or extrapolation based on the data in Table 3.
Table 3 The reproducibility limit of method 1(R)

	Mass fraction %
	Reproducibility limit (R) %

	0.73
	0.11

	1.43
	0.14

	3.00
	0.37

	Note: The reproducibility limit (R) is 2.8×SR ,where SR is the standard deviation of reproducibility.


5 EDTA Titration Method (method 2)

5.1 Principle

The samples are fused with sodium hydroxide, sodium carbonate and sodium peroxide, and extracted by hot water. The sodium carbonate washing solution is used to wash and filter out aluminum, silicon, phosphorus, etc. After filtration, the precipitate and filter paper are dissolved with hydrochloric acid and hydrogen peroxide. The filter paper is crushed into pulp. The solution and the pulp are bottled for dry filtration. In the presence of ammonium chloride, adding ammonia water to the filtrate, rare earth and iron are separated by ammonium hydroxide precipitation. In a potassium hydroxide medium with a pH of approximately 13, triethanolamine and potassium sulfate are used to mask interfering elements such as aluminum and barium. Using calcium carboxylic acid indicator as the indicator, and the calcium oxide content is determined by titration with an EDTA standard solution.
5.2 Reagents and materials 

Unless otherwise stated, only reagents confirmed to be of analytical grade or above and secondary water complying with GB/T 6682 are used in the analysis. All of the liquid reagents shall be stored in plastic bottles, certified standard solutions are preferred.

5.2.1 Sodium hydroxide.
5.2.2 Anhydrous sodium carbonate.  
5.2.3 Sodium peroxide.
5.2.4 Ammonium chloride.

5.2.5 Hydroxylamine hydrochloride.

5.2.6 Hydrogen peroxide [mass fraction 
[image: image5.wmf]w

(H₂O₂)≥30%].

5.2.7 Ammonia water (1+1).

5.2.8 Hydrochloric acid (1+1).

5.2.9 Triethanolamine (1+1).

5.2.10 Extract solution: 100 mL of water, add 2 g of sodium carbonate and 10 mL of triethanolamine solution(5.2.9).

5.2.11 Potassium sulfate solution (20 g/L).

5.2.12 Potassium hydroxide solution (400 g/L).

5.2.13 Hexamethylenetetramine solution (200 g/L).

5.2.14 Magnesium sulfate solution (50 g/L).

5.2.15 Ammonium chloride solution (20 g/L): Weigh 20 g of ammonium chloride in a conical flask and add water to 1 L to dissolve it. Then add 20 mL of ammonia water (5.2.7) to adjust the pH value of the solution to approximately 9. Boil it for use.

5.2.16 Sodium carbonate solution (10g/L).

5.2.17 Zinc standard solution [c(Zn) = 0.01000 mol/L]: Weigh 0.1625 g of pure metallic zinc (purity > 99.99%, with the surface oxide layer removed) into a 250 mL beaker, add 10 mL of water and 10 mL of hydrochloric acid (5.2.8), heat at low temperature until complete dissolution. The solution is transferred to a 250 mL volumetric flask, add 5 mL of hydrochloric acid (5.2.8), dilute it to the mark with water, and mix well.

5.2.18 Disodium ethylenediaminetetraacetate (EDTA) standard titration solution [c(EDTA) ≈ 0.010 mol/L]. Preparation and calibration are carried out according to the following requirements:

a) Preparation: Weigh 3.72 g of ethylenediaminetetraacetic acid disodium (EDTA) into a 250 mL beaker. Dissolve it with a small amount of water, transfer it to a 1000 mL volumetric flask, dilute it with water to the mark, and mix well.

b) Calibration: Transfer 15.00 mL of zinc standard solution (5.2.17) into a 250 mL conical flask, add 50 mL of water, adjust the pH of the solution to 5.0 - 5.5 using hydrochloric acid (5.2.8) or ammonia water (5.2.7), add 5 mL of hexamethylenetetramine (5.2.13), 2 drops of diphenol xylenol orange indicator (5.2.20), and titrate with EDTA standard titration solution until the solution changes from purple-red to bright yellow, indicating the endpoint. The number of parallel samples for calibration and the relative difference of the calibration results shall be shall be performed in accordance with GB/T 601.

The actual concentration c of the EDTA standard titration solution (5.2.18) is calculated according to formula (2):


[image: image6.wmf]4

3

0

×

=

V

V

c

c

…………………………………… (2)

Where,

C0 is the concentration of the zinc standard solution (5.2.17), in moles per liter (mol/L);
V3 is the volume of zinc standard solution (5.2.17) taken separately, in milliliter (mL);
V4 is the volume of EDTA standard titration solution consumed to titrate the zinc standard solution (5.2.18), in milliliter (mL).

5.2.19 Calcium carboxylic acid indicator: Calcium carboxylic acid and sodium chloride are ground together to form a solid powder in a ratio of 1:100.

5.2.20 Xylenol orange indicator (1g/L).

5.2.21 Malachite green indicator: Take 0.2 g of malachite green and dissolve in 100 mL of glacial acetic acid.

5.3 Sample

5.3.1 The particle size of the sample should not exceed 0.074 mm.

5.3.2 The sample is dried in 105 ℃ to 110℃ for 2 hours, then placed in a desiccator and cooled to room temperature.
5.4 Test Procedures

5.4.1 Test portion

Weigh the sample (5.3) according to Table 4, accurate to 0.0001 g.

Table 4 Weighing sample and transfer volume of the solution

	 Mass fraction of calcium oxide

%
	Weighing sample

g
	Transfer volume of sample solution
mL

	2.00~10.00
	0.30
	50

	>10.00~20.00
	0.30
	25


5.4.2 Parallel test

Take parallel test portions for determination and calculate the mean result.

5.4.3 Blank test 

Carry out a blank test along with the sample.

5.4.4 Preparation for analytical test solution

5.4.4.1 Transfer the sample (5.5.1) to a 30 mL nickel crucible that contains 2.0 g of dehydrated sodium hydroxide (5.2.1) and 2.0 g of anhydrous sodium carbonate (5.2.2). Cover it with 2.0 g of sodium peroxide (5.2.3). Heat it on an electric furnace to remove the moisture, then cover the crucible and place it in a 750℃ muffle furnace to melt until it turns to a bright red color and maintain for 10 minutes (take it out and shake once during the process). Then remove it and cool it.

5.4.4.2 Place the crucible and its cover in a 300 mL beaker, and add 100 mL of extract solution (5.2.10). After the intense reaction stops, rinse the crucible and its outer wall with water, add 2 mL of hydrochloric acid (5.2.8) to wash the crucible, then wash it with water and remove the crucible and its cover. Control the volume to approximately 150 mL. Boil the solution for 2 minutes, then cool slightly. Filter it using medium-speed quantitative filter paper, wash the beaker twice or three times with hot sodium carbonate washing solution (5.2.16), and wash the precipitate five or six times. Put the precipitate together with the filter paper back into the original beaker, add 20 mL of hydrochloric acid (5.2.8) and 1 mL of hydrogen peroxide (5.2.6), heat to dissolve the precipitate, destroy the filter paper, and transfer the solution along with the pulp into a 100 mL volumetric flask, dilute to the mark with water, mix it well, and dry filter.

5.4.4.3 Transfer the filtrate to a 300 mL beaker as shown in Table 4. Add 50 mL of water and then add 2 g of ammonium chloride to the beaker (5.2.4). Heat it, and add ammonia water (5.2.7) until the precipitate forms and then add an excess of 20 mL. Add 1 mL of hydrogen peroxide (5.2.6), remove it after boiling. Filter it with medium-speed filter paper into a 300 mL beaker. Wash the beaker twice to three times with the hot ammonium chloride solution (5.2.15), and wash the precipitate seven to eight times. Control the volume of the filtrate to approximately 175 mL. If it is insufficient, add the washing solution to make up to 175 mL.

5.4.5 Determination

The determination of calcium oxide content can be carried out by choosing any of the following steps.

Add 0.5 g of hydroxylamine hydrochloride (5.2.5) and 5 mL of triethanolamine solution (5.2.9) to the filtrate. Then add 20 mL of potassium hydroxide solution (5.2.12), add 5 mL of potassium sulfate solution (5.2.11) to the beaker, and let it keep quiescence for 5 minutes. Add 0.1 g of calcium carboxylic acid indicator (5.2.19), and conduct the titration using EDTA standard titration solution (5.2.18). When the color of the solution becomes faint and approaches the end point of the titration, add 1 mL of magnesium sulfate solution (5.2.14). Continue the titration, and the end point is reached when the color of the solution changes from purple to blue. 

Add 5 mL of potassium hydroxide solution (5.2.12) to the filtrate, then keep boiling for 8 minutes and subsequently cool to room temperature. Add 5 mL of triethanolamine solution (5.2.9) and 5 drops of malachite green indicator (5.2.21) to the filtrate. Adjust the solution to colorless using potassium hydroxide solution (5.2.12) and add an excess of 5 mL. Then add 0.5 g of hydroxylamine hydrochloride (5.2.5) and 5 mL of potassium sulfate solution (5.2.11), and let it keep quiescence for 5 minutes. Add 0.1 g of calcium carboxylic acid indicator (5.2.19), and conduct the titration using EDTA standard titration solution (5.2.18). When the color of the solution becomes faint and approaches the end point of the titration, add 1 mL of magnesium sulfate solution (5.2.14). Continue the titration, and the end point has been reached when the color of the solution changes from magenta to blue.

5.5 Test data processing

The calcium oxide content is calculated as a mass fraction [image: image7.wmf]w

(CaO) using to formula (3): 


[image: image8.wmf]%

100

10

08

56

-

c

a

8

2

7

3

6

5

´

´

´

´

´

´

=

-

V

m

V

.

)

V

V

(

O

C

）

（

w

…………………………………… (3)
Where,

c is the the concentration of the EDTA standard titration solution (5.2.18), in moles per liter (mol/L);
V5 is the consumed volume of the EDTA standard titration solution (5.2.18) during the titration of the test solution, in milliliter (mL);
V6 is the consumed volume of EDTA standard solution (5.2.18) during the titration of the blank test solution, in milliliter (mL);
56.08 is the molar mass of CaO, in grams per mole (g/mol);
V7 is the total volume of test solution, in milliliter (mL);
m2 is the mass of test material, in gram (g);
V8 is the volume of test solution taken separately, in milliliter (mL).
The absolute difference between two parallel determinations should not exceed the repeatability limit (r) specified in Table 5. The average value is reported as the final result. Results are retained to two decimal places and rounded conforming to the provisions of GB/T 8170.

5.6 Precision

5.6.1 Precision raw data and statistical analysis

5.6.2 Repeatability

The precision data were based on the collaborative trial statistics and calculations from 7 laboratories in 2024, using 5 different levels of samples. Each laboratory independently determined the content of calcium oxide under repeatability conditions, with 11 determinations for each level. The experimental data were statistically analyzed according to GB/T 6379.2. Under repeatability conditions, the absolute difference between two independent test results should not exceed the repeatability limit (r), and the occurrence of exceeding the repeatability limit (r) should not exceed 5%. The repeatability limit (r) is determined using linear interpolation or extrapolation based on the data in Table 5.

Table 5 The repeatability limit of method 2(r)

	Mass fraction %
	Repeatability limit (r) %

	3.11 
	0.27 

	7.23
	0.30 

	10.60
	0.37

	15.44 
	0.39 

	20.46 
	0.47 

	Note: The repeatability limit (r) is 2.8×Sr ,where Sr  is the standard deviation of repeatability.


5.6.3 Reproducibility

Under reproducibility conditions, the absolute difference between two independent test results should not exceed the reproducibility limit (R), and the occurrence of exceeding the reproducibility limit (R) should not exceed 5%. The reproducibility limit (R) is determined by linear interpolation or extrapolation based on the data in Table 6.

Table 6 The reproducibility limit of method 2(R)

	Mass fraction %
	Reproducibility limit (R) %

	3.11 
	0.28

	7.23
	0.35

	10.60
	0.46

	15.44 
	0.51

	20.46 
	0.63

	Note: The reproducibility limit (R) is 2.8×SR ,where SR is the standard deviation of reproducibility.


6 Inductively Coupled Plasma Atomic Emission Spectrometry (method 3)
6.1 Principle

The samples are fused with sodium hydroxide, sodium carbonate and sodium peroxide, and extracted by hot water. The sodium carbonate washing solution is used to wash and filter out aluminum, silicon, phosphorus, etc. The precipitate and filter paper are dissolved with hydrochloric acid. The content of calcium oxide is determined at the selected wavelength by the inductively coupled plasma atomic emission spectrometer in the hydrochloric acid medium. 

6.2 Reagents and Materials

Unless otherwise stated, only reagents confirmed to be of analytical grade or above and secondary water complying with GB/T 6682 are used in the analysis. All of the liquid reagents shall be stored in plastic bottles, certified standard solutions are preferred.

6.2.1 Sodium hydroxide.
6.2.2 Anhydrous sodium carbonate. 
6.2.3 Sodium peroxide.
6.2.4 Hydrochloric acid (ρ= 1.19 g/mL).

6.2.5 Hydrogen peroxide [mass fraction 
[image: image9.wmf]）

（

2

2

O

H

w

 ≥30%].

6.2.6 Hydrochloric acid (1+1).

6.2.7 Sodium carbonate solution (10 g/L).

6.2.8 Calcium oxide standard stored solution: Weigh 0.1785 g of calcium carbonate (with a purity of ≥ 99.99%) (previously dried at 105℃-110℃until constant weight), add 10 mL water, then add 10 mL hydrochloric acid (6.2.4) for dissolution. Transfer it into a 100 mL volumetric flask, dilute the solution with water to the mark, and mix it well. Each milliliter of this solution contains 1 mg of calcium oxide.

6.2.9 Calcium oxide standard solution: Take 5.00 mL of the calcium oxide standard stored solution (6.2.8) and transfer it into a 100 mL volumetric flask. Add 2 mL of hydrochloric acid (6.2.4), then dilute the solution with water to the mark and mix it well. this solution contains 50 μg of calcium oxide per milliliter.

6.2.10 Argon: Volume fraction ≥ 99.99%.

6.3 Apparatus

6.3.1 Inductively Coupled Plasma Atomic Emission Spectrometer.

Under the optimal working conditions of the instrument, the Inductively Coupled Plasma Atomic Emission Spectrometer can be used when it meets the following requirements:

— Resolution less than 0.007 nm (at 200 nm);

— Complies with the provisions of JJG768.

6.3.2 Electronic balance: Resolution 0.1 mg.

6.4 Sample

6.4.1 The particle size of the sample should not exceed 0.074 mm.

6.4.2 The sample is dried in 105℃ to 110℃ for 2 hours, then placed in a desiccator and cooled to room temperature.
6.5 Test Procedures

6.5.1 Test portion

Weigh 0.30 g of the sample (4.4), accurate to 0.0001 g.

6.5.2 Parallel test

Take parallel test portions for determination and calculate the mean result.

6.5.3 Blank test 

Carry out a blank test along with the sample.

6.5.4 Preparation for analytical test solution

6.5.4.1 Transfer the sample (6.5.1) to a 30 mL nickel crucible that contains 2.0 g of dehydrated sodium hydroxide (6.2.1) and 2.0 g of anhydrous sodium carbonate (6.2.2). Cover it with 2.0 g of sodium peroxide (6.2.3). Heat it on an electric furnace to remove the moisture, then cover the crucible and place it in a 750℃ muffle furnace to melt until it turns to a bright red color and maintain for 10 minutes (take it out and shake once during the process). Then remove it and cool it.

6.5.4.2 Place the crucible and its cover in a 300 mL beaker, and add 100 mL of hot water to leach the melt.. After the intense reaction stops, rinse the crucible and its outer wall with water, add 2 mL of hydrochloric acid (6.2.6) to wash the crucible, then wash it with water and remove the crucible and its cover. Control the volume to approximately 150 mL. Boil the solution for 2 minutes, then cool slightly. Filter it using medium-speed quantitative filter paper, wash the beaker twice or three times with sodium carbonate washing solution (6.2.7), and wash the precipitate five or six times. Put the precipitate together with the filter paper back into the original beaker, add 10 mL of hydrochloric acid (6.2.6) and 1 mL of hydrogen peroxide (6.2.5), heat to dissolve the precipitate, destroy the filter paper, and transfer the solution along with the pulp into a 100 mL volumetric flask, dilute to the mark with water, mix it well, and dry filter.  

6.5.4.3 Transfer the test solution (6.5.4.2) in accordance with the volume in Table 7 into a 100 mL volumetric flask, add 2 mL of hydrochloric acid (6.2.4), and dilute it to the mark with water. Mix it well.

Table 7 Transfer volume of sample solution

	Mass fraction of calcium oxide %
	Transfer volume of sample solution mL

	0.50~4.00
	10.00

	>4.00~10.00
	5.00

	>10.00~20.00
	2.00


6.5.5 Preparation of standard series solutions

Accurately transfer 0 mL, 0.50 mL, 1.00 mL, 2.00 mL, 5.00 mL, 10.00 mL and 15.00 mL of Calcium oxide standard solution (6.2.9) into a series of 50 mL volumetric flasks, respectively; add 1 mL of hydrochloric acid (6.2.4) respectively, dilute them to the mark with water, and mix them well. The mass concentration of the calcium oxide series standard solutions are 0 μg/mL、0.50 μg/mL、1.00 μg/mL、2.00 μg/mL、5.00 μg/mL、10.00 μg/mL、15.00 μg/mL respectively.

6.5.6 Determination

6.5.6.1 Plotting the calibration curve
After the Inductively Coupled Plasma Atomic Emission Spectrometer operated stably, under the selected operating conditions of the instrument, the emission intensities of the analyte element in the series of standard solutions (6.5.5) were measured at wavelengths of 393.3 nm or 396.8 nm successively. The calibration curve is plotted by the instrument software, with the mass concentration of the test elements in the standard solution series as the horizontal coordinate and the signal strength of the test element as the vertical coordinate. The correlation coefficient of the standard curve concerning each element shall exceed 0.9995; otherwise, the standardization procedure shall be performed again, or the standard solution series shall be prepared again for standardization.
6.5.6.2 Determination of both analytical test solution and blank test solution

After the calibration curve (6.5.6.1) meets the determination requirements, emission intensities of the test elements in the analytical test solution and the blank test solution are measured immediately. The instrument automatically processes the data according to the calibration curve, calculates and outputs the content of calcium oxide.
6.6 Test data processing

The calcium oxide content is calculated as a mass fraction [image: image10.wmf]w

(CaO) using to formula (4): 


[image: image11.wmf]%

100

10

-

a

11

3

6

10

9

0

´

´

´

´

´

=

-

V

m

V

V

O

C

）

（

）

（

r

r

w

…………………………………… (4)
Where,

ρ is the mass concentration of calcium oxide in the analytical test solution, in micrograms per milliliter (μg/mL);
ρ0 is the mass concentration of calcium oxid in the blank test solution, in micrograms per milliliter (μg/mL);
V9 is the total volume of test solution(6.5.4.2), in milliliter (mL);
V10 is the volume of test solution to be tested, in milliliter (mL);
m3 is the mass of test material, in gram (g);
V11 is the volume of test solution taken separately, in milliliter (mL).
The absolute difference between two parallel determinations should not exceed the repeatability limit (r) specified in Table 8. The average value is reported as the final result. Results are retained to two decimal places and rounded conforming to the provisions of GB/T 8170.

6.7 Precision

6.7.1 Precision raw data and statistical analysis

The precision data were based on the collaborative trial statistics and calculations from 7 laboratories in 2024, using 5 different levels of samples. Each laboratory independently determined the content of calcium oxide under repeatability conditions, with 11 determinations for each level. The experimental data were statistically analyzed according to GB/T 6379.2.

6.7.2 Repeatability

Under repeatability conditions, the absolute difference between two independent test results should not exceed the repeatability limit (r), and the occurrence of exceeding the repeatability limit (r) should not exceed 5%. The repeatability limit (r) is determined using linear interpolation or extrapolation based on the data in Table 8.

Table 8 The repeatability limit of method 3(r)
	Mass fraction %
	Repeatability limit (r) %

	0.74
	0.07

	1.47
	0.14

	4.69
	0.26

	10.50
	0.36

	20.42
	0.50

	Note: The repeatability limit (r) is 2.8×Sr ,where Sr  is the standard deviation of repeatability.


6.7.3 Reproducibility

Under reproducibility conditions, the absolute difference between two independent test results should not exceed the reproducibility limit (R), and the occurrence of exceeding the reproducibility limit (R) should not exceed 5%. The reproducibility limit (R) is determined by linear interpolation or extrapolation based on the data in Table 9.

Table 9 The reproducibility limit of method 3(R)
	Mass fraction %
	Reproducibility limit (R) %

	0.74
	0.10

	1.47
	0.18

	4.69
	0.33

	10.50
	0.44

	20.42
	0.60

	Note: The reproducibility limit (R) is 2.8×SR ,where SR is the standard deviation of reproducibility.



Issued by: � HYPERLINK "javascript:;" �State Administration for Market Regulation� of the People’s Republic of China


Standardization Administration of the People’s Republic of China








Implementation date: 2026-02-01 





Issue date: 2025-08-01





Chemical analysis methods for rare earth concentrates—


Part 3: Determination of calcium oxide content





稀土精矿化学分析方法


第3部分：氧化钙含量的测定


(English translation)




















National Standard of the People's Republic of China





ICS 77.120.99


CCS H 14











PAGE  
2

_1234567893.unknown

_1234567895.unknown

_1234567896.unknown

_1234567897.unknown

_1234567894.unknown

_1234567891.unknown

_1234567892.unknown

_1234567890.unknown

